OZONE SONDE CONDITIONING AND PRE-FLIGHT PREP

3-7 Days Prior to Launch

Date _______________________	Ozone Unit Number _________________Initials __________

Initial Pump Current 		___________ mA (~< 100 mA)

Pump Head Pressure 		___________ psi  (~> 10 psi = 50 cm (20 in) Hg)

Pump Suction 			___________ inches Hg (~< 50 cm (20 in) Hg)

Condition with High Ozone for 30 min  (tubing, pump, cathode chamber)			 [      ]

Condition with Zero Ozone (filter) for 5 min							[     ]

Fill with Cathode Solution (3 cc).								[     ]

Wait 2 min. Fill with Anode Solution (1.5 cc)						[     ]

Condition with Zero Ozone for 10 min							[     ]

Run on moderate ozone (5 (A) for 10 min							[     ]

Ozone free air for 10 min. Check that O3 decreases by 70-80% in 1 min				[     ]

Background current may be 0.2 - 0.5 (A at this time. It will decrease in the next days

Add 2 ml cathode solution. Short ECC-cell Leads. Store in plastic bag in cool dark place		[     ]

1 - 24 hours prior to flight

Date 				___________		____________		____________

Initials 				___________		____________		____________

Rinse Cathode and Anode cells		[    ]			[    ]			[    ]

Replace Cathode (3 cc)			[    ]			[    ]			[    ]

Replace Anode (1.5 cc) 			[    ]			[    ]			[    ]

Attach cell leads	to GCA			[    ]			[    ]			[    ]

Ozone free air on pump intake for 5 min.	[    ]			[    ]			[    ]

Pump motor current (~<100mA)	___________		____________		___________mA

Background (IB1) /~time to bkg	___________		____________		___________(A

Replace solutions if background is ~(0.10 (A (~0.35 mPa).

Moderate ozone (~5 (A) for 10 minutes	[    ]			[    ]			[    ]

Switch to ozone free air and record ((A) at: Current should decrease by 80% in first min.

0.0	  	0.5		1		3		5		10 min (IB2)

_________	_________	_________	_________	_________	_________ (1)

_________	_________	_________	_________	_________	_________ (2)

_________	_________	_________	_________	_________	_________ (3)

Flow rate (ml min-1) measured on exhaust port of cathode cell

__________		___________		___________

Vaisala attached, ozonesonde leads to TMAX board with correct polarity (blue up) Vaisala #____________

Insert heater, ozonesonde, batteries, make all connections, and

Check telemetry 				[    ]			[    ]			[    ]

Record P/T/RH	 	__________		___________		___________

Ozone Background in lab:	___________		____________		____________ hex/mPa

Background before release (IB3):__________	____________		____________ hex/mPa

Surface ozone before release (after 5-10 min):__________	____________	____________ hex/mPa

Ozone pump temp:	___________		____________		____________ (C

Twist heater leads together

Straighten Antenna

~ Surface winds at launch:	___________		____________		____________ knots

Flight Number: ____________	Flight Date: ___________		Vaisala #____________